Fig. S1. Chemical structure of HA.

uCombined=ustandard2+uweighing2+urecovery2+ucurve2

Formula S1. Equation for Combined Uncertainty Calculation.

Table S1. Suitability parameters of the chromatographic system.

| Capacity Factor | Resolution | HETP | Tailing Factor | Asymmetry Factor |
| --- | --- | --- | --- | --- |
| 7.71 | 2.8 | 0.06 | 1.1 | 0.7 |

* Values represent the average parameter measurements obtained from calibration analysis.

(a)

(b)

(c)

(d)

(e)

Fig. S2. Chromatograms of various samples: (a) Blank (aqueous medium), (b) Standard solution (10 µg/mL HA), (c) Blank cream sample, (d) Cream sample spiked with 20 µg/mL HA, (e) Real cream sample.

Table S2. Robustness of the method.

| Condition | Value | Recovery % | RSD % |
| --- | --- | --- | --- |
| Flow rate mL min-1 | 0.8 1.0 | 95.90 96.88 | 0.79 1.10 |
| Mobile phase composition (Acetonitrile:Buffer solution) | 75:25 85:15 | 93.17 95.06 | 0.46 0.95 |
| Column temperature (°C) | 25 35 | 100.40 101.27 | 0.24 0.36 |

Note: n=3 for all QC sample levels.

Table S3. Combined and expanded uncertainty results.

|  |  | Uncertainty (U) in % |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Analyte | UStandard | UCalibration | URecovery |  | URepeatability UCombined UExpanded |
| HA | 0.316 | 1.204 | 0.318 |  | 2.108 2.005 3.963 |